## 2.1 Requirements Trace Matrix
Template TMP-000024 Rev. B
2.1 Requirments Traceability Matrix
Map of the Product Design Input Requirements to Design Outputs, Design Verifications and Validations
DID ID | Design Input Requirement | Design Outputs |  | Design Verification or Validation Reports
|  | Test Plan ID | Test Case | Test Cycle | Test Results Part # | Release Cycle | Pass / Fail | Defect ID
PRSXXX | PRS Description |  | Test Case Name/Description | 1.5 |  | 1.5.54 | Pass
PRSXXX |  |  | Test Case Name/Description |  |  |  | Fail | DE3756
|  | Don’t need to fill Test Plan ID |  | These two values should be taken from application |  | These two values should be taken from application
PRS176108 | The software shall ensure the on-line help is pertinent to the applications loaded. |  | XRSCOMOH 105 001 Verify Online Help Is Specific To The Application That Is Loaded |  |  |  | FAIL | 21119, DE3576, Defect, PRS176108
PRS7282 | The systems shall provide a tutorial "Getting Started" section in the online help, similar to previous Quick reference guides. |  | XRSCOMOH 105 002 Verify Online Help Includes A Getting Started Tutorial |  |  |  | FAIL | 17017, PRS7282
PRS7283 | The systems shall provide generalized Image Quality troubleshooting for supported devices via the On-line help. |  | XRSCOMOH 105 003 Verify Image Quality Troubleshooting Via The Online Help |  |  |  | FAIL | 18204, PRS7283
PRS6275 | The software shall have the ability to display the time in one of the two formats indicated below. |  | XRSCOMLC 104 001 Verify Date And Time In Multiple Formats |  |  |  | FAIL | 14091, PRS6275
PRS6262 | Verify that the a privileged user shall be able to change the displayed GUI language to any of the supported languages. |  | XRSCOMLC 104 002 Verify User Interface For Languages |  |  |  | PASS | 14101, CIInvestigation, PRS6262
PRS6275 | The software shall have the ability to display the time in one of the two formats indicated below. | Comes from System | XRSCOMLC 104 001 Verify Date And Time In Multiple Formats | Comes from System | Comes from System | Comes from System | FAIL | 14091, PRS6275
PRS6262 | Verify that the a privileged user shall be able to change the displayed GUI language to any of the supported languages. | Comes from System | XRSCOMLC 104 002 Verify User Interface For Languages | Comes from System | Comes from System | Comes from System | PASS | 14101, CIInvestigation, PRS6262
PRS6275 | The software shall have the ability to display the time in one of the two formats indicated below. | Comes from System | XRSCOMLC 104 001 Verify Date And Time In Multiple Formats | Comes from System | Comes from System | Comes from System | FAIL | 14091, PRS6275
PRS6262 | Verify that the a privileged user shall be able to change the displayed GUI language to any of the supported languages. | Comes from System | XRSCOMLC 104 002 Verify User Interface For Languages | Comes from System | Comes from System | Comes from System | PASS | 14101, CIInvestigation, PRS6262
Note on B4: add source of design input if from CAPA, Complaint, or Product Hazard Risk